DICOM Correction Proposal

| STATUS | Letter Ballot |
| --- | --- |
| Date of Last Update | 2024/06/03 |
| Person Assigned | Björn Nolte |
| Submitter Name | Till Hoenig, till.hoenig@siemens-healthineers.com |
| Submission Date | 2023/03/21 |

Correction Number CP-2295

Log Summary: KVP attribute in CT Image Module may also be filled in multi-energy case

Name of Standard PS3.3

Rationale for Correction: With introduction of Sup188 a constraint was added to the description of (Type2) KVP attribute (0018,0060) in CT Image Module: “Shall be empty if this Attribute is present in Multi-energy CT Acquisition Sequence (0018,9362).” All other (Type 3) attributes in CT Image Module have a similar but more relaxed constraint, adding “and the value of this Attribute is not the same in all Items of the Multi-energy CT Acquisition Sequence (0018,9362).” So even if the KVP value is identical for all CT Acquisition Sequences within that image it is currently not allowed to have this value also in CT Image Module. As there are several multi-energy use cases where KVP may be or is the same value (Twin Beam, Photon-Counting, Dual-Layer-Detectors) this CP also relaxes the constraint for KVP, making it consistent with the other ME-relevant attributes on the same level. This change might also re-enable display of the KVP value (at least for above mentioned cases) on viewers which have not yet been updated to Sup188 yet. The change is seen as non-breaking.

Correction Wording:

Table C.8-3. CT Image Module Attributes

| Attribute Name | Tag | Type | Attribute Description |
| --- | --- | --- | --- |
| … |  |  |  |
| KVP | (0018,0060) | 2 | Peak kilo voltage output of the X-Ray generator used. Shall be empty if this Attribute is present in Multi-energy CT Acquisition Sequence (0018,9362) and the value of this Attribute is not the same in all Items of the Multi-energy CT Acquisition Sequence (0018,9362). Note In the context of a Multi-Energy acquisition the concerned energy spectrum needs to be considered on multiple layers: - Energy spectrum emitted by the source, see KVP values in Table C.8-3 and in Table C.8-27. - Energy spectrum effectively consumed by the detector, see Multi-energy CT X-Ray Detector Macro in Table C.8.2.2-3. - Nominal energy level associated with the image – if applicable - see Monoenergetic Energy Equivalent (0018,937C) in Multi-energy CT Characteristics Macro in Table C.8.15.3.12-1. In case there is only one KVP value for a Multi-energy acquisition, for examples Twin Beam, Photon-Counting, Dual-Layer-Detectors), the effective energies contributing to the image might be derived from the Multi-energy CT X-Ray Detector Macro in Table C.8.2.2-3. |
| … |  |  |  |