DICOM Correction Proposal

| STATUS | Letter Ballot |
| --- | --- |
| Date of Last Update | 2024/06/03 |
| Person Assigned | steven.nichols@gehealthcare.com |
| Submitter Name | steven.nichols@gehealthcare.com |
| Submission Date | 2023/08/23 |

Correction Number CP-2349

Log Summary: Relax Multi-energy CT Rescale Type constraint

Name of Standard PS3.3

Rationale for Correction: cp2189 asserted that ORIGINAL multi-energy types may include: - Objective Image Family: VMI, ELECTRON_DENSITY, EFF_ATOMIC_NUM - Material Quantification Family: MAT_SPECIFIC, MAT_REMOVED - Material Visualization Family: MAT_MODIFIED The CT Pixel Value Transformation Macro requires HU as the Rescale Type (0028,1054) for ORIGINAL images. The same conditions for Rescale Intercept (0028,1052)​ within the CT Image Module (Table C.8-3) should apply.

Correction Wording:

Modify Table C.8-126. CT Pixel Value Transformation Macro Attributes as follows:

Table C.8-126. CT Pixel Value Transformation Macro Attributes

| Attribute Name | Tag | Type | Attribute Description |
| --- | --- | --- | --- |
| Pixel Value Transformation Sequence | (0028,9145) | 1 | Contains the Attributes involved in the transformation of stored pixel values. Only a single Item shall be included in this Sequence. |
| >Rescale Intercept | (0028,1052) | 1 | The value b in relationship between stored values (SV) and the output units. Output units = m*SV + b. |
| >Rescale Slope | (0028,1053) | 1 | m in the equation specified by Rescale Intercept (0028,1052). |
| >Rescale Type | (0028,1054) | 1 | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). See Section C.11.1.1.2 for further explanation. If Frame Type (0008,9007) Value 1 of this frame is ORIGINAL and Frame Type (0008,9007) Value 3 is not LOCALIZER and Multi-energy CT Acquisition (0018,9361) is either absent or NO, the value shall be HU (Hounsfield Units). |

Editorial Note: CT Image Module is included for reference

Table C.8-3. CT Image Module Attributes

| Attribute Name | Tag | Type | Attribute Description |
| --- | --- | --- | --- |
| … |  |  |  |
| Rescale Intercept | (0028,1052) | 1 | The value b in relationship between stored values (SV) and the output units. Output units = m*SV+b If Image Type (0008,0008) Value 1 is ORIGINAL and Value 3 is not LOCALIZER, and Multi-energy CT Acquisition (0018,9361) is either absent or NO, output units shall be Hounsfield Units (HU). |
| Rescale Slope | (0028,1053) | 1 | m in the equation specified in Rescale Intercept (0028,1052). |
| Rescale Type | (0028,1054) | 1C | Specifies the output units of Rescale Slope (0028,1053) and Rescale Intercept (0028,1052). See Section C.11.1.1.2 for Defined Terms and further explanation. Required if the Rescale Type is not HU (Hounsfield Units), or Multi-energy CT Acquisition (0018,9361) is YES. May be present otherwise. |