DICOM Correction Proposal

| STATUS | Letter Ballot |
| --- | --- |
| Date of Last Update | 2024-06-03 |
| Person Assigned | David Clunie <dclunie@dclunie.com> |
| Submitter Name | Jörg Riesmeier <dicom@jriesmeier.com> |
| Submission Date | 2023-10-13 |

Correction Number CP-2364

Log Summary: Update outdated Note on Type of VOI LUT Sequence and Window Center

Name of Standard PS3.3

Rationale for Correction: A “Note” in the section about “General Requirements for Window Center and Window Width” still refers to the outdated information that the VOI LUT Sequence (0028,3010) and Window Center (0028,1050) Attributes are optional (Type 3). This information should either be updated or removed. The latter is proposed. It is also proposed to remove further inconsistencies in this section, which presumably originate from incomplete revisions in the past.

Correction Wording:

For reference PS3.3 Section C.11.2

C.11.2 VOI LUT Module

...

Table C.11-2b. VOI LUT Macro Attributes

| Attribute Name | Tag | Type | Attribute Description |
| --- | --- | --- | --- |
| VOI LUT Sequence | (0028,3010) | 1C | Defines a Sequence of VOI LUTs. One or more Items shall be included in this Sequence. Required if Window Center (0028,1050) is not present. May be present otherwise. |
| … |  |  |  |
| Window Center | (0028,1050) | 1C | Window Center for display. See Section C.11.2.1.2 for further explanation. Required if VOI LUT Sequence (0028,3010) is not present. May be present otherwise. |
| Window Width | (0028,1051) | 1C | Window Width for display. See Section C.11.2.1.2 for further explanation. Required if Window Center (0028,1050) is present. |
| … |  |  |  |

Change PS3.3 Section C.11.2.1.2.2

C.11.2.1.2.2 General Requirements for Window Center and Window Width

The Window Center (0028,1050), Window Width (0028,1051) and VOI LUT Function (0028,1056) Attributes shall be used only for Images with Photometric Interpretation (0028,0004) values of MONOCHROME1 and MONOCHROME2. They have no meaning for other Images.

If multiple values are present in the Window Center (0028,1050) and Window Width (0028,1051) Attributes, both Attributes shall have the same number of values and shall be considered as pairs. Multiple values indicate that multiple alternative views may be presented.

If any VOI LUT Table, specified by the VOI LUT Sequence (0028,3010) Attribute, is included by an Image, a pair of Window WidthCenter (0028,1050) and Window CenterWidth (0028,1051) values, or the VOI LUT Table, but not both, may be applied to the Image for display. Inclusion of both indicates that multiple alternative views may be presented.

If multiple Items are present in VOI LUT Sequence (0028,3010), only one may be applied to the Image for display. Multiple Items indicate that multiple alternative views may be presented.

If the VOI LUT Module is defined in an IOD and if neither a VOI LUT Sequence (0028,3010) nor a Window WidthCenter (0028,1050) and Window CenterWidth (0028,1051) are present, then the VOI LUT stage of the grayscale pipeline (described in PS3.4 Section N.2) is defined to be an identity transformation.

Note

1.	This requirement is specified so that IODs that define a particular output space for the grayscale pipeline, such as P-Values, are not in an undefined state when no VOI LUT Sequence or Window Width and Window Center are present.

2.	Despite the Type 3 requirement for Though the VOI LUT Module, VOI LUT Sequence (0028,3010), and Window Center (0028,1050) and Window Width (0028,1051) Attributes may be optional in a particular IOD, implementations that render images are expected to implement and apply these transformations when they are present in the image, unless overridden by the user, a presentation state, or a hanging protocol, and to allow the user to select which transformation to apply when multiple transformations are present.