DICOM Correction Proposal

| STATUS | Letter Ballot |
| --- | --- |
| Date of Last Update | 2024-06-03 |
| Person Assigned | David Clunie <dclunie@dclunie.com> |
| Submitter Name | Jörg Riesmeier <dicom@jriesmeier.com> |
| Submission Date | 2024-01-02 |

Correction Number CP-2385

Log Summary: Fix inconsistent names of Transfer Syntaxes

Name of Standard PS3.11

Rationale for Correction: Throughout DICOM PS3.11, the names of the Transfer Syntaxes are not consistent with the official names associated with the Transfer Syntax UIDs according to PS3.6. This results in a wrong use of terms, e.g., in DICOM Conformance Statements (at least when based on the “old” DCS template according to DICOM PS3.2-2022d Annex A). In this context, there are also other inconsistencies that are proposed to be fixed with this CP, e.g., in relation to the difference between Application Profile and profile class. Editorial change: The spelling of defined terms such as “Application Profile” should be harmonized in this Part of the DICOM Standard, i.e., written with capital initial letters. The same is true for terms like “SOP Instance”, “Transfer Syntax”, “File-set Creator”, “File-set Updater”, “File-set Reader”, etc.

Correction Wording:

Change PS3.11 Section A.3.1

A.3.1 SOP Classes and Transfer Syntaxes

This Application Profile is based on the Media Storage Service Class (see PS3.4).

SOP Classes and corresponding Transfer Syntaxes supported by this Application Profile are specified in the Table A.3-1.

Table A.3-1. STD-XABC-CD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| X-Ray Angiographic Image | 1.2.840.10008.5.1.4.1.1.12.1 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Mandatory | Mandatory | Optional |

Note

1.	This aApplication pProfile does not allow the use of the X-Ray Angiographic Bi-Plane Image IObjectD. Bi-plane acquisitions must therefore be transferred as two single plane SOP iInstances. A future Application Profile that permits X-Ray Angiographic Bi-Plane Image Object transfer is under development.

2.	This Application Profile includes only the X-Ray Angiographic Image SOP Instances. It does not include Standalone Curve, Modality LUT, VOI LUT, or Overlay SOP Instances.

Change PS3.11 Section B.3.1

B.3.1 SOP Classes and Transfer Syntaxes

This Application Profile Class is based on the Media Storage Service Class (see PS3.4).

SOP Classes and corresponding Transfer Syntaxes supported by this Application Profile are specified in Table B.3-1.

Table B.3-1. STD-XA1K-CD and STD-XA1K-DVD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement (see Note 1) |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| X-Ray Angiographic Image | 1.2.840.10008.5.1.4.1.1.12.1 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Mandatory | Mandatory | Optional |
| X-Ray Angiographic Image | 1.2.840.10008.5.1.4.1.1.12.1 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) 1.2.840.10008.1.2.4.50 | Optional for DVD; Disallowed for CD | Mandatory for DVD; Disallowed for CD | Undefined for DVD; Disallowed for CD |
| X-Ray Angiographic Image | 1.2.840.10008.5.1.4.1.1.12.1 | JPEG Extended (Process 2 & 4): Default Transfer Syntax for Lossy JPEG 12 Bit Image Compression (Process 4 only) 1.2.840.10008.1.2.4.51 | Optional for DVD; Disallowed for CD | Mandatory for DVD; Disallowed for CD | Undefined for DVD; Disallowed for CD |
| Secondary Capture Image Storage | 1.2.840.10008.5.1.4.1.1.7 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory | Optional |
| Grayscale Softcopy Presentation State Storage | 1.2.840.10008.5.1.4.1.1.11.1 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional | Optional |

Note

1.	The FSU requirement is not defined for STD-XA1K-DVD profile.

2.	The Standalone Overlay, Standalone Curve and Detached Patient mManagement SOP Classes were formerly defined in these profiles, but have been retired. The Grayscale Softcopy Presentation State Storage SOP Class has been added as the preferred mechanism for conveying annotations.

Change PS3.11 Section C.3.1

C.3.1 AbstractSOP Classes and Transfer Syntaxes

Application Profiles in this class, STD-US, shall support the appropriate Information Object Definitions (IOD) and Transfer Syntaxes for the Media Storage SOP Class in the following table. In the role of FS-Updater or FS-Creator, the application can choose one of the three possible tTransfer Syntaxes to create an IOD SOP Instance. In the role of FS-Reader, an application shall support all tTransfer Syntaxes defined for the respective STD-US aApplication pProfile.

Table C.3-1. Ultrasound SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name | Transfer Syntax UID |
| --- | --- | --- | --- |
| DICOM Media Storage Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed | 1.2.840.10008.1.2.1 (see Section 8.6 in PS3.10) |
| Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 | Explicit VR Little Endian Uncompressed | 1.2.840.10008.1.2.1 |
| Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 | RLE Lossless Image Compression | 1.2.840.10008.1.2.5 |
| Ultrasound Image Storage | 1.2.840.10008.5.1.4.1.1.6.1 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) | 1.2.840.10008.1.2.4.50 |
| Ultrasound Multi-frame Image Storage | 1.2.840.10008.5.1.4.1.1.3.1 | Explicit VR Little Endian Uncompressed | 1.2.840.10008.1.2.1 |
| Ultrasound Multi-frame Image Storage | 1.2.840.10008.5.1.4.1.1.3.1 | RLE Lossless Image Compression | 1.2.840.10008.1.2.5 |
| Ultrasound Multi-frame Image Storage | 1.2.840.10008.5.1.4.1.1.3.1 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) | 1.2.840.10008.1.2.4.50 |

C.3.1.1 Ultrasound Single and Multi-frame Pixel Formats Supported

The Application Profiles of the STD-US application profileclass requires that all ultrasound image objects only be stored using the values described in PS3.3the US Image Module in PS3.3 and the specializations used for the Ultrasound Single and Multi-Frame IODs.

In the role of FS-Updater or FS-Creator the application can choose any of the supported Photometric Interpretations described in PS3.3the US Image Module in PS3.3 to create an IOD SOP Instance. In the role of FS-Reader, an application shall support all Photometric Interpretations described in PS3.3the US Image Module in PS3.3.

Table C.3-2 describes restrictions on the use of various Transfer Syntaxes with the supported Photometric Interpretations for both single and multi-frame images.

Table C.3-2. Defined Photometric Interpretation and Transfer Syntax Pairs

| Photometric Interpretation Value | Transfer Syntax Name | Transfer Syntax UID |
| --- | --- | --- |
| MONOCHROME2 | UncompressedExplicit VR Little Endian RLE Lossless Image Compression | 1.2.840.10008.1.2.1 1.2.840.10008.1.2.5 |
| RGB | UncompressedExplicit VR Little Endian RLE Lossless Image Compression | 1.2.840.10008.1.2.1 1.2.840.10008.1.2.5 |
| PALETTE COLOR | UncompressedExplicit VR Little Endian RLE Lossless Image Compression | 1.2.840.10008.1.2.1 1.2.840.10008.1.2.5 |
| YBR_FULL | RLE Lossless Image Compression | 1.2.840.10008.1.2.5 |
| YBR_FULL_422 | UncompressedExplicit VR Little Endian JPEG LossyBaseline (Process 1) | 1.2.840.10008.1.2.1 1.2.840.10008.1.2.4.50 |

Change PS3.11 Section C.3.3.2

C.3.3.2 File Component IDs

Note

File Component IDs should be created using a random number filename to minimize File Component ID collisions as described in PS3.12. The FS-Updater should check the existence of a Component ID prior to creating that ID. Should an ID collision occur, the FS-Updater should try another ID.

Change PS3.11 Section D.3.1

D.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class (see PS3.4).

Table D.3-1. STD-GEN STD-GEN-CD, STD-GEN-DVD-RAM, STD-GEN-SEC-CD, STD-GEN-SEC-DVD-RAM, STD-GEN-BD, STD-GEN-SEC-BD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| Composite Image & Stand-alone Storage Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Defined in Conformance Statement | Defined in Conformance Statement | Optional |

The SOP Classes and corresponding Transfer Syntax supported by thiese Application Profiles are specified in the Table D.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section E.3.1

E.3.1 SOP Classes and Transfer Syntaxes

These Application Profiles are based on the Media Storage Service Class (see PS3.4).

SOP Classes and corresponding Transfer Syntaxes supported by these Application Profiles are specified in the Table E.3-1.

Table E.3-1. STD-CTMR SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement (see Note 1) |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| CT Image | 1.2.840.10008.5.1.4.1.1.2 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Optional | Mandatory | Optional |
| CT Image | 1.2.840.10008.5.1.4.1.1.2 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory | Optional |
| MR Image | 1.2.840.10008.5.1.4.1.1.4 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Optional | Mandatory | Optional |
| MR Image | 1.2.840.10008.5.1.4.1.1.4 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory | Optional |
| Secondary Capture Image (Grayscale) | 1.2.840.10008.5.1.4.1.1.7 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Optional | Mandatory | Optional |
| Secondary Capture Image (Grayscale) | 1.2.840.10008.5.1.4.1.1.7 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory | Optional |
| Secondary Capture Image (Palette Color) | 1.2.840.10008.5.1.4.1.1.7 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Optional | Optional | Optional |
| Secondary Capture Image (Palette Color) | 1.2.840.10008.5.1.4.1.1.7 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional | Optional |
| Grayscale Softcopy Presentation State | 1.2.840.10008.5.1.4.1.1.11.1 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional | Optional |
| X-Ray Radiation Dose SR | 1.2.840.10008.5.1.4.1.1.88.67 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional | Optional |

Note

1.	The FSU requirement is not defined for the STD-CTMR-DVD profile.

2.	The Detached Patient mManagement SOP Class was formerly defined in these profiles, but has been retired.

Change PS3.11 Section G.3.1

G.3.1 SOP Classes and Transfer Syntaxes

This Application Profile is based on the Media Storage Service Class (see PS3.4).

Table G.3-1. STD-GEN-MIME SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional |
| Composite Image & Stand-alone Storage Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Defined in Conformance Statement | Defined in Conformance Statement | Defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntaxes supported by this Application Profile are specified in the Table G.3-1. The supported Storage SOP Class(es) and Transfers Syntax(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section G.3.1

H.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class (see PS3.4).

Table H.3-1. STD-GEN-DVD and STD-GEN-SEC-DVD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) 1.2.840.10008.1.2.4.50 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Extended (Process 2 & 4): Default Transfer Syntax for Lossy JPEG 12 Bit Image Compression (Process 4 only) 1.2.840.10008.1.2.4.51 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression (Lossless Only) 1.2.840.10008.1.2.4.90 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression 1.2.840.10008.1.2.4.91 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntaxes supported by thiese Application Profiles are specified in the Table H.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section I.3.1

I.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class (see PS3.4).

Table I.3-1. STD-DVD-MPEG2-MPML and STD-DVD-SEC-MPEG2-MPML SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG2 Main Profile @/ Main Level Image Compression 1.2.840.10008.1.2.4.100 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntax supported by thiese Application Profiles are specified in the Table I.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section J.3.1

J.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class (see PS3.4).

Table J.3-1. STD-GEN-USB, STD-GEN-SEC-USB, STD-GEN-MMC, STD-GEN-SEC-MMC, STD-GEN-CF, STD-GEN-SEC-CF, STD-GEN-SD and STD-GEN-SEC-SD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) 1.2.840.10008.1.2.4.50 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Extended (Process 2 & 4): Default Transfer Syntax for Lossy JPEG 12 Bit Image Compression (Process 4 only) 1.2.840.10008.1.2.4.51 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression (Lossless Only) 1.2.840.10008.1.2.4.90 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression 1.2.840.10008.1.2.4.91 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntaxes supported by thiese Application Profiles are specified in the Table J.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section K.3.1

K.3.1 SOP Classes and Transfer Syntaxes

The Application Profile STD-DEN-CD shall support the SOP Classes and Transfer Syntaxes in the following table.

Table K.3-1. Dental AbstractSTD-DEN-CD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory |
| Digital Intra-oral X-Ray Image Storage - For Presentation | 1.2.840.10008.5.1.4.1.1.1.3 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory |
| Digital X-Ray Image Storage - For Presentation | 1.2.840.10008.5.1.4.1.1.1.1 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory |
| Basic Structured Display Storage | 1.2.840.10008.5.1.4.1.1.131 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional |
| Grayscale Softcopy Presentation State | 1.2.840.10008.5.1.4.1.1.11.1 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Optional |

Note

The Digital X-Ray Image Storage and Digital Intra-oral X-Ray Image Storage - For Presentation SOP Classes can also be used for scanned film.

A File S-set Creator (FSC) shall support at least one of the specified image storage SOP Classes.

Change PS3.11 Section L.3.1

L.3.1 STD-GEN-ZIP-MAIL and STD-GEN-SEC-ZIP-MAIL AbstractSOP Classes and Transfer Syntaxes

Applications interchanging data under the STD-GEN-ZIP-MAIL and STD-GEN-SEC-ZIP-MAIL profiles shall support the Information Object Definitions (IOD) and Transfer Syntaxes for the Media Storage SOP Class specified in Table L.3-1.

Table L.3-1. STD-GEN-ZIP-MAIL and STD-GEN-SEC-ZIP-MAIL SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory |
| Composite Image & Stand-alone Storage Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Defined in Conformance Statement | Defined in Conformance Statement | Defined in Conformance Statement |

Equipment claiming conformance to these Application Profiles shall list the subset of Media Storage SOP Classes and Transfer Syntaxes that it supports in its Conformance Statement.

Change PS3.11 Section L.4.1

L.4.1 STD-DTL-SEC-ZIP-MAIL AbstractSOP Classes and Transfer Syntaxes

Applications interchanging data under the STD-DTL-SEC-ZIP-MAIL profile shall support the Information Object Definitions (IOD) and Transfer Syntaxes for the Media Storage SOP Class specified in Table L.3-2. File S-set Creators for the STD-FTL-SEC-ZIP-MAIL shall support at least one of the optional IODs.

Table L.3-2. STD-DTL-SEC-ZIP-MAIL AbstractSOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement |
| --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory |
| Digital Intra-oral X-Ray Image Storage - For Presentation | 1.2.840.10008.5.1.4.1.1.1.3 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory |
| Digital X-Ray Image Storage - For Presentation | 1.2.840.10008.5.1.4.1.1.1.1 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Optional | Mandatory |

Change PS3.11 Section M.3.1

M.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class with the Interchange Option (see PS3.4).

Table M.3-1. STD-GEN-BD and STD-GEN-SEC-BD SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossless Non-Hierarchical, First-Order Prediction (Process 14 (s[Selection vValue 1)] 1.2.840.10008.1.2.4.70 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Lossy, Baseline Sequential with Huffman Coding (Process 1) 1.2.840.10008.1.2.4.50 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG Extended (Process 2 & 4): Default Transfer Syntax for Lossy JPEG 12 Bit Image Compression (Process 4 only) 1.2.840.10008.1.2.4.51 | Defined in Conformance Statement | Mandatory for -JPEG profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression (Lossless Only) 1.2.840.10008.1.2.4.90 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | JPEG 2000 Image Compression 1.2.840.10008.1.2.4.91 | Defined in Conformance Statement | Mandatory for -J2K profiles for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG2 Main Profile @/ Main Level 1.2.840.10008.1.2.4.100 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG2 Main Profile @/ High Level 1.2.840.10008.1.2.4.101 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG-4 AVC/H.264 High Profile / Level 4.1 1.2.840.10008.1.2.4.102 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG-4 AVC/H.264 BD-compatible High Profile / Level 4.1 1.2.840.10008.1.2.4.103 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntaxes supported by thiese Application Profiles are specified in the Table M.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.

Change PS3.11 Section N.3.1

N.3.1 SOP Classes and Transfer Syntaxes

Theise Application Profiles isare based on the Media Storage Service Class with the Interchange Option (see PS3.4).

Table N.3-1. STD-GEN-BD-MPEG4-LV42 and STD-GEN-SEC-BD-MPEG4-LV42 SOP Classes and Transfer Syntaxes

| Information Object Definition | SOP Class UID | Transfer Syntax Name and UID | FSC Requirement | FSR Requirement | FSU Requirement |
| --- | --- | --- | --- | --- | --- |
| Basic Directory | 1.2.840.10008.1.3.10 | Explicit VR Little Endian Uncompressed 1.2.840.10008.1.2.1 | Mandatory | Mandatory | Mandatory |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG-4 AVC/H.264 High Profile / Level 4.2 For 2D Video 1.2.840.10008.1.2.4.104 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG-4 AVC/H.264 High Profile / Level 4.2 For 3D Video 1.2.840.10008.1.2.4.105 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |
| Multi-frame Composite IODs for which a Media Storage SOP Class is defined in PS3.4 | See PS3.4 | MPEG-4 AVC/H.264 Stereo High Profile / Level 4.2 1.2.840.10008.1.2.4.106 | Defined in Conformance Statement | Mandatory for all SOP Classes defined in Conformance Statement | Defined in Conformance Statement |

The SOP Classes and corresponding Transfer Syntaxes supported by thiese Application Profiles are specified in the Table N.3-1. The supported Storage SOP Class(es) shall be listed in the Conformance Statement using a table of the same form.