Audience Chat from “Implementation of Digital Pathology at CHUM”- Jan 20 2023

| # | From | Chat | Answer | Comment |
| --- | --- | --- | --- | --- |
|  | Monief Eid | Does your scanner support all DICOM classes including the worklist? And how you carry out the QC of rejected slides? Last but not least does your LIS reporting support the structure reporting? Monief | Our LIS does not support level 6 synoptic reporting. We do have synoptic reports for cancer cases in free text form. | First Question was addressed in Q&A |
|  | Monief Eid | Are the scanners from one vendor or different? and if from different , do you see any difference in quality of slides scanning and how do you manage that? Monief |  | Question was addressed in Q&A |
|  | David Beyer | How long do you keep images in hot storage? 30 days? longer? Thanks |  | Question was addressed in Q&A |
|  | Monief Eid | For tele-pathology, do you view through zero foot print viewer and or web browser? and how is the quality of image remotely? | We have access to both: through zero foot print viewer and web browser. The image quality is the same as on site. |  |
|  | Josh Lewis | What technical staff assess the digital images before assigning to the pathologist? | Medical technologists | Question was addressed in Q&A |
|  | Gurdip Tamber | Are cytology samples validated for WSI? | No not yet |  |
|  | Yuri Nikolski | Is your implementation a stand alone system or is it integrated with Radiology/Cardiology imaging either via PACS or workflow orchestrator? Are the images archived? If so, where are they currently archived? (ie: separate archive, VNA etc) | We do not archive our images at the moment, except for research and teaching cases. | First Question was addressed in Q&A |
|  | Monief Eid | congratulations for such big success , amazing |  | Kudos! |
|  | Gurdip Tamber | how long each WSI is stored for? | 4 months after last modification date | Question was addressed in Q&A |
|  | Monief Eid | to Yuri question, is this part of VNA? |  |  |
|  | David Clunie | You mentioned interoperability - how is this address between scanners and storage and viewers (DICOM?), and between LIS and scanners (HL7? IHE DPIA?), and between facilities with different vendors equipment for scanning and storage. | LIS-IMS: HL7 Scanner-IMS: common access to an image folder generated by scanner | First Question was addressed in Q&A, but I think this was the question you were to follow-up upon. |
|  | Dean Jin | In scoping out storage requirements for Digital Pathology, is there any average size for a DP study that could be considered? | This depends on the slide volume of your service, the duration of storage, the type of storage (hot or cold), the use of archived images (research, teaching, diagnosis). For instance, for some research projects, there is no need to keep WSIs in their original versions that contain many layers of redundant images, thus the required storage capacity should be lower than for diagnostic purposes. | Question was addressed in Q&A, but consider further info if able. |
|  | David Clunie | Do you plan to share WSI with the regional Diagnostic Image Repository (DI-r)? |  | Question was addressed in Q&A |
|  | Damian Hitchins | Do you still keep the original slides? | Yes as per Quebec regulatory rules (10 years) |  |
|  | Damian Hitchins | Has the Quebec government adjusted the storage requirements for the original slides or is this just digitized and if you keep short term digital and still keep the original how frequently do you rescan? | See above. We rescan exceptionally very few cases. |  |
|  | David Beyer | Do you think there is interest in creating a national level WSI de-identified database (ie within Canada) ? I know in many provinces there is interest in moving digital pathology forward; and this would be incredibly useful for developing AI tools in our public system (I say this as an Alberta-based pathologist) | I can only express my keen interest for this initiative. |  |
|  | Damian Hitchins | Do you keep key slides for long term storage? So you can bring back the original. Additionally do you ever still review through the microscope original slide because you are not satisfied with a digitized slide? | Yes we started to keep 1-2 representative digital slides per new cases. Infrequently, we do have to review glass slides for polarization, special stains for very small microorganisms (Zield etc.) since the highest magnification of our scanners is 40X, or slides that cannot not be properly scanned due to technical issues (folds, uneven samples etc.) |  |
|  | Josh Lewis | Medical Laboratory Technologists or Medical Laboratory Technician? | Medical Laboratory Technologists operate the scanners, technicians help out charging and unloading slide trays at our lab. | I believe this question was addressed in Q&A |
|  | Iris Teo | For the AI algorithms (e.g. KI67, slide quality),could you provide insight into any QA/QC that is routinely performed to ensure that the algorithms perform as initially designed? | For each case, the responsible pathologist review marked nuclei to ensure that the algorithm is accurate. |  |
|  | Thomas Jay Crawford | You showed an array of different monitors. What steps are undertaken to ensure consistent display output between workstations? | Only one of the monitor is used to view digital images. We followed recommendations from DPA once, at the beginning, when setting up work stations for pathologists. During training period, a second round of adjustment is possible if there is discrepancy between the images from glass slides and digital slides. |  |
|  | Dimitrios Divaris | Can you or when would you be able to real time to do Radiology pathology correlations using medical images images and Pathology images | Radiology Pathology Correlation and Concordance has been a topic of discussions for the last decade in the Imaging Informatics domain. Presently radiologist and pathologist communicate ad hoc at times and probably not as frequent as it should be! The basic concept, is improving communications between pathologists and radiologists regarding a patient’s imaging and a patient’s path results will improve diagnosis, accuracy of diagnosis and facilitate self-learning. It is thought these flows from Rad to Path as well as Path to Rad, dependent on the timeline of the case procedure and testing. UCLA developed a system t facilitate RadPath Correlation, https://pubmed.ncbi.nlm.nih.gov/26521686/ Recently, IHE (Integrating the Healthcare Enterprise) in partnership with the College of American Pathologists (CAP) and the American College Of Radiologists (ACR) Data Institute, published a Profile on Radiology Pathology Concordance (RPC). This a Technical Framework Supplement for Trial Implementation. This publication is intended for use by vendors and solution developers to produce systems. Here is more information on this IHE initiative . In short, there may be solutions that facilitate better and easier Radiology-Pathology communications and sharing of Radiology-Pathology results in the near future. But, if we do not ask for these types of solutions from our health administrators, these vendors will not build them. | See comment from M. David Kwan |
|  | Bruno Bellotto | Just a comment.....my pathology manager says: "We take approx. 800 digital photos per year and file size ranges from 252KB to 5600KB – most are between 400 and 900KB" | A whole slide image (WSI) is not a photo. From DPA website: WHOLE SLIDE IMAGE (WSI): A digitized histopathology glass slide that has been created on a slide scanner. The digitized glass slide represents a high-resolution replica of the original glass that can then be manipulated through software to mimic microscope review and diagnosis. Also referred to as a virtual slide. Thus, a WSI is composed of multiple layers or tiles of images assembled in a single virtual slide. |  |
|  | Bruno Bellotto | so ~1MB per slide | The average is 1-1.5 GB per WSI |  |
|  | Andi Broka | Where will the cloud storage be hosted? Is it on a public cloud vendor like Amazaon, Google, etc? | Cloud storage for CHUM will likely be on a public cloud vendor |  |
|  | Monief Eid | Thank you Jason and David for this session opportunity and sharing knowledge with others ….Monief |  | Kudos! |
|  | David Beyer | Thanks! That was very informative |  | Kudos! |
|  | Wes McLachlin | Thank you! Excellent presentation |  | Kudos! |
|  | Dean Jin | Thank you. Awesome presentation! |  | Kudos! |
|  | Cody Alexander | Thank you! |  | Kudos! |
|  | Kerry Bezan | thank you very interesting |  | Kudos! |
|  | Aman Ish Goyal | Very informative. Thank you, Dr. Nguyen. Thanks David, Jason for hosting this webinar. |  | Kudos! |
|  | Marina MacPherson | Merci beaucoup! |  | Kudos! |